## Slide 1: JPE presentation
Active vibration isolation for Cold Finger Einstein Telescope
Huub Janssen
20-04-2023

## Slide 2
JPE

## Slide 3
Einstein Telescope – ET Pathfinder
How is JPE involved?
Design of cryogenic vibration isolator.
Test evaluation.
Einstein Telescope

## Slide 4
Einstein Telescope

## Slide 5
ET Pathfinder Universiteit Maastricht

## Slide 6
Mirror suspension

## Slide 7
Cold Finger
Cold finger

## Slide 8
Proven track record on accurate positioning in vacuum and cryogenic environment.
In house development of cryogenic positioners with nanometric resolution.
Vibration challenges from customers in scientific research.
How is JPE involved?

## Slide 9
Requirements
  XYZ isolation of Cold Finger
  Compact volume (250 mm diameter and 200 in height)
  Low payload mass
  Resonance frequency about 1 Hz (low pass filter)
Design of cryogenic vibration isolator

## Slide 10
Mechatronic concept
The mechanical design is based on the following key items:
Completely frictionless guides of payload mass and sensor mass by introduction of flexure elements.
Magnetic gravity compensation of sensor mass.
Force actuation by means of cryogenic compatible voice coil actuator.
Displacement measurement by means of optical interferometer.
Gravity compensation of payload mass by mechanical spring.

## Slide 11
Mechanical layout

## Slide 12
Mechanical layout

## Slide 13
Transmissibility = xp/xf [-] (goal: << 1)
Compliance = xp/Fd [m/N] (goal: << 1); Fd=xd * Cbraid [N]
Test evaluation
Cold Finger
Vibration Isolator
payload table
Floor
Cooler
Mirror
xf
xd
Xp
C braid

## Slide 14
Transmissibility and Compliance
Theoretical expectations
Active isolation starts at 0.33 Hz while passive isolation starts at 10 Hz
Active compliance starts at 0.33 Hz while passive compliance starts at 10 Hz

## Slide 15
Transmissibility
In passive mode: resonances of both the suspended inertial- and payload mass are present at 4 Hz and 8 Hz respectively.
In active mode: damped resonance at 0.3 Hz remaining.
Measured results

## Slide 16
Vibration rejection ratio – active isolation
Inverse of transmission ratio

## Slide 17
Compliance. (predicted)
Measured results

## Slide 19
Realized a cryogenic compatible compact active isolator with a relatively small mass that is able to significantly attenuate floor vibrations and reduce disturbance forces on the payload above 0.3 [Hz]
Great improvement found of the passive to active isolation frequency:
  Passive resonance frequencies are 4 and 8 Hz with very little amount of damping
  Active effective resonance frequency around 0.3 Hz with a large amount of damping.
Conclusions

## Slide 20
Realization of XYZ vibration isolator
Next steps